BIOSAFETY CABINETS

Biosafety cabinets are designed for controlling aerosols and micro-particles associated with managing potentially toxic or infectious biological material in laboratories in activities such as agitation, centrifugation, pipetting, and opening of pressurized containers. Safety cabinets have been designed to protect the user, the environment and the sample (biological material) manipulated using appropriate ventilation conditions. They are also known as laminar flow cabinets and/or biosafety cabinets.

The biological safety cabinets are used for the following:

1. To protect the worker from risks associated with the management of potentially infectious biological material.

2. To protect the sample being analyzed from becoming contaminated.

3. To protect the environment.

- Biosafety cabinets are only used for certain risk group organisms and for processes that might result in aerosol formation.

- These cabinets are provided with HEPA-filters that decontaminate the air moving out of the cabinet.

- Biosafety cabinets might be confused with the laminar hood as both of these pieces of equipment work as enclosed workspaces. But, laminar hood only provides protection to the sample and not to the personnel and the environment, whereas biosafety cabinets protect all three.

- The use of biosafety cabinets or other such physical containment is not required in the biosafety level 1, but depending on the risk assessment, some processes might require such containment.

- BSCs are an essential part of biosafety as they minimize the formation of aerosol, protecting the environment, the pathogen, and the laboratory personnel.

- Besides, most BSCs also function to sterilize biological materials that are kept inside the cabinets.

OPERATION PRINCIPLE

The biological safety cabinet is a chamber generally constructed of steel. It has a front glass window of adjustable height, a ventilation system with an electrical motor, a ventilator and a set of ducts which while functioning, generate a negative pressure condition inside the cabinet. This forces the air to flow from inside the cabinet through the front opening to generate a curtain of air protecting the operator. Internally, the air is conducted through a series of grids and ducts to be finally treated in HEPA 1 filters. Depending on the design of the cabinet, the air is recycled inside the laboratory or extracted and renewed in diverse proportions. The air flow, which in Class II cabinets moves from the filter towards the work surface, is laminar.

Biosafety Cabinet Classes

Biosafety cabinets are classified into three classes (Class I, Class II and Class III) by the U.S. Centers for Disease Control and Prevention (CDC), each with specific performance characteristics and applications.

Class I and II Biosafety cabinets are used for biosafety risk levels I and II but, when used correctly in conjunction with useful microbiological techniques, these provide an effective containment system for safe manipulation of moderate and high-risk microorganisms.

Class III BSCs are most suitable for work with hazardous agents that require Biosafety Risk Level 3 or 4.

Risk level 1 group is composed of biological agents very unlikely to cause sickness in healthy humans or animals. (No individual and community risk).

2. Risk level 2 group is composed of pathogens which cause sickness in humans or animals but unlikely to be dangerous to laboratory workers, the community, domestic animals or the environment under normal circumstances. Those exposed in the laboratory rarely become seriously ill. There are preventive measures and effective treatment available and the risk of dissemination is limited.

3. Risk level 3 group is composed of pathogens which usually cause serious sicknesses to human beings and animals and produce a serious economic impact. However, infection by casual contact by one by one individual to another is not common. The sicknesses these produce are treatable by antimicrobial or anti- parasitic agents.

4. Risk level 4 group is composed of pathogens which usually produce very serious sicknesses in human beings or animals, frequently without treatments available. These agents are easily spread from one individual to another or from animal to human being or vice versa, directly or indirectly or by casual contact.

Class I

- Class I is the most basic biosafety cabinet that provides protection to the environment and the laboratory personnel.

- It doesn’t, however, provide protection to the product as the unsterilized room air is drawn over the work surface.

- Class I biosafety cabinets are typically used to either enclose specific equipment like centrifuges or for procedures like aerating cultures that might potentially generate aerosols.

- Biosafety cabinets of this class are either ducted (connected to the building exhaust system) or unducted (recirculating filtered exhaust back into the laboratory).

- In the Class I BSC, the room air is drawn in through the opening that also allows the entry of the operator’s arm during work.

- The air inside the cabinet then takes in the aerosol particles that may have been generated and moves it away from the operator towards the HEPA filter.

- The air moving out of the cabinet is thus, sterilized via the HEPA filters before its discharge to the environment.

- In this way, the cabinets protect the operator and the environment from the aerosol but not the sample.

Class II

- BSC-Class II cabinets provide protection of the samples and the environment since makeup air is also HEPA-filtered.

- The principle of operation of Class II cabinets involves a fan mounted in the top of the cabinet that draws a curtain of sterile air over the workstation where the biological products are being handled.

- The air then moves underneath the work station and back up to the top of the cabinet before passing through the HEPA filters.

- The exhaust that moves out of the facility consists of air being drawn into the front of the cabinet underneath the work surface.

- The air drawn in acts as a barrier against the potentially contaminated air coming back out to the operator.

- Class II BSCs are further divided into five types depending on the exhaust system and the mechanism of work (recirculation of the exhaust air); Type A1, Type A2, Type B1, Type B2, and Type C1.

### a. Type A1

- The type A1 cabinets have a minimum inflow velocity of 75ft/min where the contaminated air divided just above the work station and mixes with the inflow air.

- The mixed air is then drawn through a duct network so that it reaches the back of the cabinet.

- After this, air might be either recirculated after passing through the HEPA filters or exhausted out of the cabinet, also through a HEPA filter.

- This type of cabinet is not as widely used as it is not safe to work with hazardous chemical substances.

### b. Type A2

- The type A2 cabinets have a minimum inflow velocity of 100 ft/min.

- In Class II Type A2 BSC, air enters the chamber through the front aperture, which provides operator protection.

- The inflow air mixes with the down flow air (from the top of the cabinet) and enters the front intake grille and then passes over the workstation where the air splits.

- Approximately 60% to 70% of the contaminated air is recycled and pushed back into the workstation in the chamber through the down flow HEPA filter, while the remaining 30% to 40% is exhausted through the exhaust HEPA filter.

- However, if hazardous, volatile chemicals are to be used within the cabinet along with the microbiological work, exhaust must be released into the atmosphere through the direct duct system.

- Because of the chances of the release of hazardous chemicals into the environment, A2 type cabinets are also not extensively used.

### c. Type B1

- Type B cabinets are different from Type A cabinets as they use single-pass airflow to control the flow of hazardous vapors.

- Type B1 cabinets divide the airflow so that the contaminated air is directed towards the exhaust system while the air between the operator and the workstation mixes with the inflow and is recirculated.

- The exhaust air dispersed out of the facility should be passed through the HEPA filters to provide protection to the environment.

- These cabinets have a dedicated duct system which allows the release of the contaminated air out of the facility.

- In the case of Type B1 cabinets, 40% of the air is recirculated, whereas the remaining 60% is exhausted out of the facility.

### d. Type B2

- For a Type B2 BSC, like in Type A cabinets, the air is drawn in from the front opening creating an air barrier that protects the operator.

- Air is also drawn in from an opening at the top of the cabinet that supplies the down flow of air in the cabinet.

- The air then passes through a HEPA filter, where 100% of the air is exhausted through a dedicated duct system with an exhaust fan motor. The air moving out of the facility is thus sterilized before its release into the atmosphere.

- The advantage of this system is the removal of toxic vapors that are generated in the cabinet with no recirculation within the BSC.

- All of the contaminated airflow (100%) in a Type B2 cabinet is externally exhausted which means the air drawn into the cabinet is 100% exhausted into the atmosphere.

- As a result, none of the air drawn into the B2 from either inflow or down flow is recycled within the airflow system.

- Because none of the air is recirculated, these cabinets are the best to be used for tasks involving the release of chemical vapors.

- Type B2 cabinets, however, are expensive, and their use is limited to toxicology laboratories where protection against hazardous chemicals is imperative.

### e. Type C1

- Type C1 cabinets are similar to Type B cabinets in their working mechanism, but these are designed to reduce operating costs and add flexibility to the laboratories.

- These cabinets work by using the single-pass airflow system where the cabinets move the air by mixing it with the down flow air separated into columns for recirculation.

- The air above the workstation is drawn with a second fan which moves the contaminated air out through the exhaust system with a HEPA filter.

- In this way, the cabinets provide protection to the environment, the operator, and the workstation or the biological material (sample).

- Type C cabinets are different from Type A cabinets as they use a single-pass airflow mechanism where the air is not circulated.

- They differ from Type B cabinets in that they don’t require a dedicated ducted exhaust system, can work for an extended duration to increase operator protection in the case of exhaust failure, and can even run without the exhaust at all.

Class III

- Class III cabinets are leak-tight, totally enclosed but ventilated cabinets, where all air that either enters or leaves through the facility pass through a HEPA filter.

- The cabinets are provided with rubber gloves that are attached to the system to be used during operations in the cabinet. This is why these cabinets are also termed ‘glove boxes’.

- The cabinet even has a transfer chamber that facilitates the sterilization of materials before they leave the glove box.

- Even though the gloves restrict the hand movement of the operator inside the cabinet, it prevents direct contact between the operator and the samples.

- The exhaust air is treated with double HEPA filters or HEPA filters in combination with incineration.

- These cabinets can be used for all four Biosafety risk levels (1, 2, 3, and 4). They are the most important cabinets for the manipulation of biological materials in the Biosafety level 4.

- These cabinets are mostly custom-built for specific laboratories with lab equipment built inside the chamber.

- All of the structural and design features of these cabinets provide maximum protection to the operator, the environment, and the sample against the high-risk group 4 pathogenic organisms.

USE OF THE SAFETY CABINET

1. Plan the work to be done in the biological safety cabinet in advance. Determine what procedure and equipment will be used. Coordinate the time of the cabinet’s use with the other laboratory professionals in order to avoid interruption or undesired traffic while it is in use.

2. Turn on the cabinet. Turn off the UV lamp if lit. Turn on the fluorescent light lamp and the cabinet’s ventilator. Verify that the grids in front and behind are free of obstructions. Prepare the work area. Allow the cabinet to function for at least 15 minutes.

3. Wash hands and forearms with germicidal soap. Put on the personal protective apparel: coat/overall with long sleeves and adjustable cuff s, protective eyeglasses and mask if the work requires it. Prepare the interior surfaces of the cabinet applying 70% ethanol or a suitable disinfectant. After this, let the air flow through.

4. Only load and install the materials and equipment required for the test or manipulation. Distinguish between the clean areas and dirty areas. Place the material in such a way that the clean materials do not mix or cross used or dirty materials or impede the circulation of the internal air

through the front and back grids. Place a biosafety bag for disposing waste materials, a container with disinfectant for the pipettes and a container for storing sharps. Avoid locating very large objects near one another. Upon finalizing the placing of the materials, the flow of air must be allowed to sweep through the cabinet for approximately 3 to 5 minutes in order to eliminate any particle produced or freed during the loading of materials and equipment.

5. Initiate activities. Slowly introduce hands into the work area. Carry on the processes and tasks in a methodical and careful manner (from the clean areas to the potentially contaminated areas). Keep the materials at least 10 cm behind the front grid. Try to perform the most risky and contaminating activities towards the back of the cabinet’s work area. Avoid the use of open flames of lighters since this breaks the laminar flow pattern and may burn the filter.

Avoid removing hands from the work area until all procedures are

Accomplished and the potentially dangerous materials are disposed of in the biosafety bag or in the pipette and sharp containers established procedure. Wash hands with a lot of water and soap. Turn off the ventilator, the fluorescent lamp, close the front opening and turn on the

Ultraviolet light.

Specialized maintenance

The following are some procedures to be done according to the manufacturer’s technical service manuals by a specialized contractor.

1. Annual certification in accordance with Regulation NSF 49 outlines.

2. Motor change. Generally, it uses maintenance-free sealed rollers and function by induction through frequency control. This motor does not have brushes.

3. Replacing ventilators when its faults.

4. Replacing the HEPA filter. The replacement frequency depends on the use of the cabinet and the system of environmental control installed in the laboratory. If there is a good control of dust, the

filters could last many years.

5. Repair of the electronic control system: flow control alarms, position of the window, velocity controls.

6. Repair/cleaning of the flow regulator valves, bell type adjustment fittings.

The certification process for BSC TEST:

1. Air tightness test. This is done on the exterior surfaces. Determine if joints, seals, penetration and soldering are free from leaks.

2. HEPA filters leak tests. Determines the integrity of the supply and extraction of HEPA filters, their location and mounted frames.

3. Temperature increase test. Determines the maximum temperature increase in the cabinet when the ventilator and lights are operating.

4. Noise test. Determines the level of noise produced by the cabinet.

5. Luminous intensity test. Determines the luminous intensity on the cabinet’s work surface.

6. Vibrations test. Determine how much vibration there is in the cabinet when it is functioning.

7. Protection test to personnel, to the product and cross contamination biological tests. The test determines if aerosols are contained in the cabinet, if external contaminants reach the work table area and if aerosols are reduced by the cabinet.

8. Stability test. Determines if the cabinet has structural stability. Analyzes the resistance to shaking, to distortion by means of applied force, to deflection of the work surface subjected to load and

Resistance to the tilting of the work surface due to heavy loading conditions.

9. Vertical flow velocity test. Determines the velocity of the air moved vertically towards the work surface.

10. Entry flow velocity test. Determines the velocity at which the fl ow enters the cabinet through the front opening and the cabinet’s extraction volume.

11. Smoke test. Determines if the flow of air along the entire perimeter of the front opening advances towards the cabinet, and if the vertical flow moving towards the bottom does not show dead points or flow backs on the work surface.

12. Drainage escape test. Defines the contention capacity for spills below the work surface.

13. Motor/ventilator system functioning test. Determines if the system provides the necessary static pressure.

14. Electric system test. Determines if there are potential risks of electrical discharges. Measures the escaping currents, the polarity, the functioning of the ground defect protection system and the ground circuit resistance.

BY MR OKUKU